## Slide 1
Development of Novel Semiconductor Detectors In Xiantan University
新型半导体粒子、光子探测器的研发
Zheng Li 李正
Xiangtan University, Xiangtan,湘潭大学 Hunan 411105, China
2018.09

## Slide 2
OUTLINE
研究背景
一
Large area Si drift detectors (SDD)
2
Summary
4
High resistivity Si Detectors
1
3D-Terench-Electrode Detectors
3

## Slide 3
Appication of high resistivity Si Detectors
半导体探测器的应用
Neutrino
Higgs Factory
Nuclear Reactor
Nuclear test monitoring
Nuclear warning
X-ray and X-ray Pulsar-based Navigation（XPNAV）
Hand-hold monitor
Light sources
Dirty bomb monotor

## Slide 4
氧化 OXIDATION
双面光刻
DOUBLESIDED-LITHO
匀胶 PHOTORESIS
百级超净室 (Class-100 Clearnroom，220m2)
Si detector processing in Xiangtan

## Slide 5
Operation Principles
Thin entrance window
(thin dead layer) needed for soft X-rays, and ions
3) e’s and h’s move and separate
in the field to generate signal
d : thickness is hundreds
of m’s, up to 1mm
Width and length are
up to 100 mm
1. High Resistivity Si Detector

## Slide 6
P (pitch)
d : up to 1 mm
W : up to 100 mm
L : up to 100 mm
P : down to 50 m
Total detector volume
V = LWd is active!
d (wafer thickness)
W (detector width)
L (strip length)
p+
n-type Si
n+
Single-side Si Strip Detectors
h
-
-
-
-
-
-
-
-
Typical Detector Geometries
1D-Position sensitivity

## Slide 7
1d position sensitivity, single-sided processing
P (pitch)
d : up to 1 mm
W : up to 100 mm
L : up to 100 mm
P : down to 50 m
Total detector volume
V = LWd is active!
d (wafer thickness)
W (detector width)
L (strip length)
p+
n-type Si
n+
-
-
-
-
-
-
-
-
Single-side Si Strip Detectors
Q
1D-Position sensitivity

## Slide 8
n-type
 4 kΩ-cm
V  +100 volts
X (µm)
0
XP (Position of the sensing strip)
d = 100 to 600 µm
Y (µm)
YP (Position of the sensing strip)
Single-metal technology, double-sided litho
MIP
2D-Position sensitivity
Double-side Si Strip Detectors

## Slide 9
n-type
 4 kΩ-cm
V  +100 volts
X (µm)
0
XP (Position of the sensing strip)
d = 100 to 600 µm
Y (µm)
YP (Position of the sensing strip)
Single-metal technology, double-sided litho
Double-side Si Strip Detectors
2D-Position sensitivity

## Slide 10
n+
p+
p+
n-type
Si
  Operating principle
Single-metal technology, double-sided litho
    mm’s to cm’s
    250 to
    500 m
    Anode (n+)
    p+
    p+
h or particle
h or particle
Si Drift Detectors （SDD）

## Slide 11
n+
p+
p+
n-type
Si
  Operating principle
Single-metal technology, double-sided litho
Si Drift Detectors （SDD）
    mm’s to cm’s
    250 to
    500 m
    Anode (n+)
Q
    p+
    p+

## Slide 12: large area detector with a small collection anode
Widely used for X-ray spectroscopy with its low noise (low capacitance (60 fF)).
Excellent energy resolution
Large area detectors
Short shaping: high count rate applications
For photon science and space (NASA lunar mission)
Single-metal technology, double-sided litho
    Large area SDD array
Si Drift Detectors （SDD）
    5 mm
a)
b)
c)

## Slide 13
  Operating principle
h or particle
  Conventional
  small area SDD
  with constant
  backside bias
Large area Si Drift Detectors D = 2 cm

## Slide 14
Ultra Large area Si Drift Detectors
Φ= 5 cm (R=2.5cm)， Near constant drift field
Potential distributions
Field distributions
Hexagonal Cell

## Slide 15
Ultra Large area Si Drift Detectors
Φ= 5 cm (R=2.5cm), Near linear drift field
Potential distributions
Field distributions
Hexagonal Cell

## Slide 16
Vout=250V
VB,out=170V
VE1=10V
VB,E1=83V
250
200
150
100
50
0
0.3
0.5
0.9
0.7
1.0
0.025
0.05
x (cm)
r (cm)
0.1
ø(r,x) (V)
Phi(r,x) data
m=1, n=4
Potential distributions
Vout
VBout
VB
VE1

## Slide 17
Vout=250V
VB,out=170V
VE1=10V
VB,E1=83V
250
200
150
100
50
0
0.3
0.5
0.9
0.7
1.0
0.01
0.02
0.03
0.04
0.05
x (cm)
r (cm)
Phi(r,x) data
m=1, n=4
ø(r,x) (V)
Potential distributions
Vout
VBout
VB
VE1

## Slide 18
Electric field distribution
Vout just right
VB just right
Anode
Drift channel is within and pointing to
anode

## Slide 19
Electric field distribution
Vout just right
VB just right
Drift channel is within and pointing to
anode
Anode

## Slide 20
Electric field distribution
Vout too large
Drift channel is out
of detector body

## Slide 21
Electric field distribution
VB too large
Drift channel is not pointing to anode
Anode

## Slide 22
Electric field distribution
VB too large
Vout too large
Drift channel is out of
detector body and not
pointing to anode
Anode

## Slide 23
SDD Design (7 layears/Masks, douple-sided)
314 mm2
600 mm2
SDD Mask Set Design 掩膜版试验设计

## Slide 24
Fabricated SDD (Top side) SDD 晶片正面
Fabricated SDD (Bottom side) SDD 晶片反面
314 mm2
600 mm2
SDD Mask Set Design掩膜版试验设计
大面积硅漂移探测器）
Large area Si Drift Detectors D = 2 cm

## Slide 25
Top-side 314mm2
大面积硅漂移探测器）（正面）
Large area Si Drift Detectors D = 2 cm

## Slide 26
（SDD设计细节）
Bottom-side 314mm2
大面积硅漂移探测器）（反面）
Large area Si Drift Detectors D = 2 cm

## Slide 27
探测器芯片测试技术
Test system (Multi-I, Multi-V, Multi-Probe Station)
测试过程及操作示范 (多电流，多电压源， 多探针）
C - V 测试数据
I - V 测试数据
SDD Test results

## Slide 28
d
W
Depletion depth W is decoupled from the thickness d
The New 3D-Trench-Electrode Detectors
V
Full depletion voltage is small and independent on d
Pixel cells are isolated from each other --- Independent Coaxial Detector Array
(ICDA)
Trench electrode
Collection electrode (column)
“New BNL 3D-Trench electrode Si detectors for radiation hard detectors for sLHC and for X-ray applications”
Z. Li, Nucl. Instr. and Meth. A 658 (2011) 90–97 doi:10.1016/j.nima.2011.05.003

## Slide 29
d : up to 1 mm
R : up to 500 m
3D-Trench-Electrode Si Detectors
3D-processing technology
Single-metal technology, double-sided litho
Operation Principle
d (wafer thickness)
n-type Si
n+ column
p+-trench
R
-
-
-
-
-
-
-
-
h
Q

## Slide 30
d : up to 1 mm
R : up to 500 m
3D-processing technology
Single-metal technology, double-sided litho
d (wafer thickness)
n-type Si
n+ column
p+-trench
MIP
R
-
-
-
-
-
-
-
-
-
-
Q
The New 3D-Trench-Electrode Detectors
Operation Principle

## Slide 31
150 μm
300 μm
400 μm
p+ trench
Trench and column 10 μm wide
n-type bulk
5x1011 cm-3
p-spray implant the back
Hexagonal geometry (n-type column) for photon science
( R=150 µm)
n+ column
Two electrodes
3D-Trench-ORJ
Full 3D simulations
a)
b)

## Slide 32
Electric Field (V/cm)
Electron conc. (cm-3)
Hole conc. (cm-3)
2nd cut (2D 1D)
5x1011
0
1200
1000
800
600
200
400
1st cut (3D 2D)
Full depletion
already at Vbias=4V
(d = 400 μm)
Actural 3D detecter
(D=160 μm)
a)
b)
c)
d)
e)
f)

## Slide 33
Simulated capacitance: C = 99 fF
Theoretical capacitance: C = 97.8 fF
l=270 μm R/r=10
Similar to SDD, it has very small capacitance, good for X-ray applications
Capacitance

## Slide 34
a
b
2D cross section at z=135 μm. a) electric field for non-irradiated detetor, V=1 V; b) electric field at the fluence of 1x1016 neq/cm2, V=100 V.
Electric field simulation
Fully depleted at
V = 1V
V = 100V

## Slide 35
10 V
60 V
90 V
30 V
Hole concentration
distribution at
1x1016 neq/cm2
Full depletion
voltage is
about 90 V
For non-irradiaed
it is <1 V!
Hole concentration (depletion trend)

## Slide 36
Collected charge for the non-irradiated detector
Very fast <1 ns

## Slide 37
Collected charge for the irradiated detector
(c)
(d)
Radiation hard

## Slide 38
Fabricated detectors: One example:
Hexangular type, R = 80 m, 49x54 array
Trenches
250 um depth
Design
Etched trenches
Actual photos

## Slide 39
Fabricated detectors: on 4" wafer
100 mm

## Slide 40
R = 500 m
1 mm

## Slide 41
Hex, R = 80 mc
160 μm

## Slide 42
Square, L = 212 m
212 μm

## Slide 43
Leakage current test

## Slide 44
20V
60V
100V
Fast
Full
Hex 02 100 um from column
Aug 8th
Laser scan results, NYU 8/2014

## Slide 45
60V
100V
140V
Hex 01
Laser scan
results, NYU
8/2014
60V

## Slide 46
Laser Study of 49x54 3D-Trench Electrode Detector Array
2017 results
49x54 3D-Trench Electrode Detector Array

## Slide 47
1
2
3
4
5
6
7
Area of Detail
Laser Study of 49x54 3D-Trench Electrode Detector Array
2017 results
160 m

## Slide 48
View of Wire-Bonds
Laser Study of 49x54 3D-Trench Electrode Detector Array
2017 results

## Slide 49
Laser Settings
Pulse Voltage: 1.56V
Pulse Width: 20nsec
Averaging: 128
Step Size: 0.01mm
Filtering Settings
High Pass
  C=22nF
  R=510Ω
Low Pass
  R=50Ω
  C=1nF
  Detector Voltage
  Bias -100V
  Guard Ring floating
Laser Study of 49x54 3D-Trench Electrode Detector Array
2017 results

## Slide 50
Chart: 
| -48.9 | -48.89 | -48.88 | -48.87 | -48.86 | -48.85 | -48.84 | -48.83 | -48.82 | -48.81 | -48.8 | -48.79 | -48.78 | -48.77 | -48.76 | -48.75 | -48.74 | -48.73 | -48.72 | -48.71 | -48.7 | -48.69 | -48.68 | -48.67 | -48.66 | -48.65 | -48.64 | -48.63 | -48.62 | -48.61 | -48.6 | -48.59 | -48.58 | -48.57 | -48.56 | -48.55 | -48.54 | -48.53 | -48.52 | -48.51 | -48.5 | -48.49 | -48.48 | -48.47 | -48.46 | -48.45 | -48.44 | -48.43 | -48.42 | -48.41 | -48.4 | -48.39 | (blank)
-76.99 | 0.15365 | 0.16172 | 0.27052 | 0.17156 | 0.16365 | 0.26557 | 0.20443 | 0.38584 | 0.29929 | 0.35534 | 0.3758 | 0.34347 | 0.50252 | 0.4391 | 0.34074 | 0.24332 | 0.25794 | 0.18743 | 0.26267 | 0.33893 | 0.19805 | 0.29023 | 0.38505 | 0.32587 | 0.37459 | 0.39241 | 0.41759 | 0.4058 | 0.31434 | 0.26555 | 0.21229 | 0.22655 | 0.17299 | 0.23058 | 0.7134 | 0.22143 | 0.47665 | 0.24258 | 0.39161 | 0.29575 | 0.34458 | 0.61302 | 0.55775 | 0.28748 | 0.43481 | 0.30287 | 0.44586 | 0.23934 | 0.43577 | 0.18288 | 0.14244 | 0.14007
-76.98 | 0.33853 | 0.37339 | 0.41432 | 0.3764 | 0.19679 | 0.25515 | 0.45309 | 0.2777 | 0.42107 | 0.41518 | 0.36352 | 0.41351 | 0.36873 | 0.39139 | 0.3409 | 0.44984 | 0.52023 | 0.45677 | 0.29713 | 0.26447 | 0.33517 | 0.27575 | 0.52514 | 0.43086 | 0.56239 | 0.35013 | 0.42584 | 0.33096 | 0.4977 | 0.74949 | 0.27816 | 0.22501 | 0.62921 | 0.2609 | 0.28805 | 0.34788 | 0.29745 | 0.66526 | 0.58384 | 0.38471 | 0.57261 | 0.34374 | 0.30808 | 0.31776 | 0.34642 | 0.35007 | 0.27127 | 0.31687 | 0.28925 | 0.36727 | 0.33302
-76.97 | 0.27915 | 0.18531 | 0.17391 | 0.20936 | 0.277 | 0.4662 | 0.35001 | 0.27721 | 0.4061 | 0.32413 | 0.58718 | 0.33838 | 0.48284 | 0.44124 | 0.53315 | 0.56193 | 0.37616 | 0.38876 | 0.48422 | 0.2851 | 0.40507 | 0.37477 | 0.3266 | 0.37593 | 0.43279 | 0.39821 | 0.5016 | 0.30529 | 0.48337 | 0.3605 | 0.96702 | 0.3406 | 0.31895 | 0.32436 | 0.47744 | 0.25553 | 0.27802 | 0.32755 | 0.38719 | 0.44921 | 0.44696 | 0.43456 | 0.3882 | 0.38903 | 0.45796 | 0.2537 | 0.48349 | 0.32686 | 0.22075 | 0.1937 | 0.16345
-76.96 | 0.1881 | 0.22103 | 0.23203 | 0.27091 | 0.47996 | 0.29349 | 0.62467 | 0.46983 | 0.44702 | 0.28474 | 0.33379 | 0.41205 | 0.43409 | 0.41224 | 0.45 | 0.42107 | 0.48825 | 0.49513 | 0.38066 | 0.46348 | 0.46305 | 0.61638 | 0.42422 | 0.33589 | 0.43394 | 0.34583 | 0.3732 | 0.39478 | 0.30606 | 0.39658 | 0.43696 | 0.45509 | 0.54627 | 0.4847 | 0.43223 | 0.38007 | 0.3913 | 0.43041 | 0.43409 | 0.60971 | 0.55345 | 0.49712 | 0.58546 | 0.46294 | 0.45488 | 0.5318 | 0.3905 | 0.3855 | 0.22082 | 0.25251 | 0.46733
-76.95 | 0.31602 | 0.5435 | 0.36036 | 0.26044 | 0.35399 | 0.26408 | 0.26213 | 0.33808 | 0.35134 | 0.46805 | 0.52089 | 0.43253 | 0.52749 | 0.40597 | 0.5339 | 0.66509 | 0.4346 | 0.54363 | 0.73732 | 0.65549 | 0.39462 | 0.41856 | 0.4619 | 0.5411 | 0.43353 | 0.49122 | 0.46083 | 0.70903 | 0.54891 | 0.46862 | 0.46454 | 0.37516 | 0.47072 | 0.52002 | 0.38068 | 0.36688 | 0.47656 | 0.74015 | 0.48541 | 0.55541 | 0.48971 | 0.4832 | 0.50679 | 0.47102 | 0.53051 | 0.48162 | 0.41399 | 0.33064 | 0.60435 | 0.18997 | 0.19244
-76.94 | 0.20166 | 0.13814 | 0.28063 | 0.60776 | 0.44192 | 0.32842 | 0.58377 | 0.34488 | 0.39858 | 0.58143 | 0.51514 | 0.53591 | 0.62667 | 0.56084 | 0.60697 | 0.69833 | 0.76086 | 0.62988 | 0.84716 | 0.62257 | 0.46671 | 0.50933 | 0.48678 | 0.71859 | 0.59322 | 0.60741 | 0.59992 | 0.48329 | 0.52487 | 0.59233 | 0.50063 | 0.53384 | 0.55055 | 0.74494 | 0.5958 | 0.71532 | 0.50078 | 0.48148 | 0.57371 | 0.58661 | 0.57538 | 0.55673 | 0.52667 | 0.49883 | 0.56014 | 0.42036 | 0.6554 | 0.30663 | 0.25947 | 0.4089 | 0.16353
-76.93 | 0.18444 | 0.24845 | 0.25995 | 0.35327 | 0.34642 | 0.34528 | 0.33181 | 0.41297 | 0.42233 | 0.48583 | 0.55893 | 0.53595 | 0.65888 | 0.62775 | 0.57181 | 0.7372 | 0.58934 | 0.65136 | 0.79339 | 0.82355 | 0.60669 | 0.57251 | 0.5839 | 0.61645 | 0.82724 | 0.59237 | 0.56458 | 0.56891 | 0.59828 | 0.74865 | 0.54032 | 0.54628 | 0.54334 | 0.6669 | 0.61719 | 0.54987 | 0.87594 | 0.58526 | 0.5875 | 0.57313 | 0.54243 | 0.54802 | 0.53236 | 0.55899 | 0.5124 | 0.46662 | 0.39752 | 0.31403 | 0.31744 | 0.27003 | 0.16811
-76.92 | 0.37573 | 0.20095 | 0.48126 | 0.24522 | 0.34149 | 0.38186 | 0.36862 | 0.41802 | 0.7388 | 0.58119 | 0.65342 | 0.64748 | 0.68748 | 0.81116 | 0.74197 | 0.71861 | 0.68697 | 0.72457 | 0.75523 | 0.71044 | 0.69063 | 0.6761 | 0.77899 | 0.76937 | 0.70331 | 0.64292 | 0.6282 | 0.86951 | 0.60599 | 0.59462 | 0.66245 | 0.62497 | 0.82944 | 0.83462 | 0.68326 | 0.55426 | 0.62633 | 0.69478 | 0.75323 | 0.64695 | 0.55929 | 0.62438 | 0.69637 | 0.50227 | 0.56107 | 0.40068 | 0.47082 | 0.30194 | 0.59073 | 0.28213 | 0.19722
-76.91 | 0.28242 | 0.35178 | 0.25808 | 0.27051 | 0.31835 | 0.38188 | 0.60289 | 0.63271 | 0.57434 | 0.69372 | 0.82383 | 0.89618 | 0.94426 | 0.83772 | 0.83224 | 0.85516 | 0.88901 | 0.66554 | 0.78337 | 0.77505 | 0.78822 | 0.79233 | 0.93638 | 0.89038 | 0.9271 | 0.83825 | 0.89972 | 0.78066 | 0.86038 | 0.87525 | 0.68785 | 0.66182 | 0.7832 | 0.67036 | 0.63075 | 0.71738 | 0.70725 | 0.72854 | 0.95747 | 1.0068 | 0.69499 | 0.64361 | 0.62296 | 0.55707 | 0.47076 | 0.39246 | 0.44455 | 0.32773 | 0.30003 | 0.24432 | 0.16466
-76.9 | 0.25663 | 0.20764 | 0.4122 | 0.28509 | 0.47721 | 0.3947 | 0.52256 | 0.47766 | 0.54053 | 0.74309 | 0.79714 | 0.96964 | 1.0834 | 0.91644 | 1.0604 | 0.9419 | 0.92678 | 0.92344 | 0.92775 | 0.88486 | 0.94144 | 0.94825 | 1.0104 | 1.0002 | 1.0895 | 0.94563 | 0.90412 | 1.1257 | 0.90789 | 0.77909 | 0.85943 | 0.74219 | 0.7629 | 0.98919 | 0.83415 | 0.72073 | 0.89826 | 0.95405 | 0.92742 | 0.90689 | 0.87624 | 0.8222 | 0.6774 | 0.91276 | 0.4788 | 0.43753 | 0.34932 | 0.56615 | 0.54592 | 0.57338 | 0.18931
-76.89 | 0.21848 | 0.22943 | 0.25624 | 0.27983 | 0.38606 | 0.43282 | 0.55001 | 0.51291 | 0.79278 | 0.85125 | 0.90378 | 1.1727 | 1.0889 | 1.1059 | 1.0204 | 1.0982 | 1.1926 | 1.2109 | 1.2464 | 1.0528 | 1.1365 | 1.2602 | 1.1414 | 1.1111 | 1.1329 | 0.9995 | 0.93365 | 0.944 | 0.95536 | 0.94024 | 1.0056 | 1.0857 | 1.0395 | 0.9019 | 0.96166 | 1.0653 | 1.0317 | 0.98765 | 1.1067 | 1.0028 | 1.2358 | 0.82442 | 0.62284 | 0.65817 | 0.51798 | 0.42958 | 0.38215 | 0.45214 | 0.27105 | 0.23947 | 0.45046
-76.88 | 0.259 | 0.25635 | 0.27738 | 0.49854 | 0.3799 | 0.5985 | 0.59361 | 0.54462 | 0.67531 | 0.73031 | 0.90557 | 1.0778 | 1.2545 | 1.2774 | 1.2453 | 1.2798 | 1.3864 | 1.4278 | 1.5503 | 1.3648 | 1.3989 | 1.272 | 1.3554 | 1.2714 | 1.1664 | 1.0395 | 1.0086 | 1.1933 | 1.0686 | 1.0683 | 1.3643 | 1.1551 | 1.2427 | 1.2711 | 1.347 | 1.2164 | 1.3702 | 1.1667 | 1.1566 | 1.1106 | 0.98674 | 0.85418 | 0.85515 | 0.74261 | 0.56859 | 0.58577 | 0.38126 | 0.368 | 0.33019 | 0.25189 | 0.24013
Pixels 1 & 2
160 m
Laser Study of 49x54 3D-Trench Electrode Detector Array
2017 results

## Slide 51
Chart: 
Pixels 3,4,5
160 m
160 m
Laser Study of 49x54 3D-Trench Electrode Detector Array
2017 results

## Slide 52
Chart: 
| -48.9 | -48.89 | -48.88 | -48.87 | -48.86 | -48.85 | -48.84 | -48.83 | -48.82 | -48.81 | -48.8 | -48.79 | -48.78 | -48.77 | -48.76 | -48.75 | -48.74 | -48.73 | -48.72 | -48.71 | -48.7 | -48.69 | -48.68 | -48.67 | -48.66 | -48.65 | -48.64 | -48.63 | -48.62 | -48.61 | -48.6 | -48.59 | -48.58 | -48.57 | -48.56 | -48.55 | -48.54 | -48.53 | -48.52 | -48.51 | -48.5 | -48.49 | -48.48 | -48.47 | -48.46 | -48.45 | -48.44 | -48.43 | -48.42 | -48.41 | -48.4 | -48.39 | (blank)
-77.19 | 0.20625 | 0.18625 | 0.22081 | 0.25932 | 0.23214 | 0.3463 | 0.27114 | 0.34593 | 0.31662 | 0.33783 | 0.41856 | 0.40343 | 0.36564 | 0.33997 | 0.4288 | 0.37613 | 0.31831 | 0.3649 | 0.34115 | 0.29943 | 0.277 | 0.31707 | 0.35846 | 0.41084 | 0.36918 | 0.41941 | 0.49788 | 0.40032 | 0.34494 | 0.4923 | 0.34537 | 0.34287 | 0.35192 | 0.33075 | 0.24995 | 0.32727 | 0.30168 | 0.24286 | 0.24116 | 0.32022 | 0.33308 | 0.32974 | 0.28923 | 0.27332 | 0.22699 | 0.31842 | 0.28411 | 0.227 | 0.25957 | 0.22433 | 0.1828 | 0.17879
-77.18 | 0.2422 | 0.28411 | 0.24802 | 0.30571 | 0.29937 | 0.30627 | 0.39429 | 0.363 | 0.31368 | 0.34967 | 0.43378 | 0.44706 | 0.43944 | 0.39107 | 0.35967 | 0.45378 | 0.4562 | 0.42653 | 0.46195 | 0.3965 | 0.4002 | 0.38336 | 0.36728 | 0.3579 | 0.33108 | 0.40699 | 0.41456 | 0.34347 | 0.45526 | 0.42484 | 0.40842 | 0.36223 | 0.38374 | 0.4319 | 0.33233 | 0.35391 | 0.35561 | 0.42247 | 0.3777 | 0.282 | 0.35414 | 0.29086 | 0.24865 | 0.25799 | 0.21669 | 0.26261 | 0.40081 | 0.30807 | 0.13883 | 0.17154 | 0.15668
-77.17 | 0.24522 | 0.25906 | 0.27635 | 0.27895 | 0.42537 | 0.33147 | 0.50025 | 0.32659 | 0.32137 | 0.37196 | 0.41003 | 0.44769 | 0.41265 | 0.50936 | 0.57603 | 0.46837 | 0.52549 | 0.47854 | 0.62248 | 0.43415 | 0.4549 | 0.36866 | 0.45722 | 0.40044 | 0.40662 | 0.50583 | 0.50287 | 0.49236 | 0.37812 | 0.38976 | 0.39787 | 0.46137 | 0.4715 | 0.47913 | 0.43662 | 0.32504 | 0.35215 | 0.27549 | 0.37127 | 0.3651 | 0.38092 | 0.32838 | 0.24442 | 0.36138 | 0.37654 | 0.3148 | 0.30656 | 0.30257 | 0.34223 | 0.16766 | 0.24238
-77.16 | 0.23678 | 0.18431 | 0.22526 | 0.32109 | 0.44387 | 0.34527 | 0.40079 | 0.41506 | 0.42971 | 0.44965 | 0.46217 | 0.6381 | 0.57754 | 0.55737 | 0.60717 | 0.5354 | 0.4955 | 0.4112 | 0.60276 | 0.74397 | 0.46223 | 0.44878 | 0.52954 | 0.49029 | 0.52336 | 0.59074 | 0.58271 | 0.58073 | 0.62112 | 0.42153 | 0.4981 | 0.45569 | 0.54125 | 0.57271 | 0.5866 | 0.47463 | 0.4148 | 0.37745 | 0.39308 | 0.39432 | 0.3006 | 0.32334 | 0.387 | 0.29435 | 0.33025 | 0.23784 | 0.1921 | 0.26632 | 0.31622 | 0.20832 | 0.14307
-77.15 | 0.18419 | 0.20177 | 0.20792 | 0.30717 | 0.48901 | 0.43056 | 0.4386 | 0.36941 | 0.52544 | 0.62228 | 0.59355 | 0.57685 | 0.6559 | 0.57658 | 0.69929 | 0.60503 | 0.59964 | 0.54696 | 0.77901 | 0.7527 | 0.59896 | 0.61826 | 0.5885 | 0.61334 | 0.67097 | 0.70097 | 0.66362 | 0.58552 | 0.60086 | 0.5548 | 0.56371 | 0.46534 | 0.47624 | 0.63475 | 0.61267 | 0.53918 | 0.47129 | 0.47301 | 0.47142 | 0.4474 | 0.32837 | 0.33562 | 0.3541 | 0.35876 | 0.30867 | 0.28254 | 0.32126 | 0.25774 | 0.36998 | 0.16166 | 0.19892
-77.14 | 0.26906 | 0.15162 | 0.19024 | 0.19567 | 0.42634 | 0.42684 | 0.45175 | 0.43563 | 0.54382 | 0.68304 | 0.62068 | 0.63243 | 0.85495 | 0.75072 | 0.75442 | 0.72799 | 0.67638 | 0.64329 | 0.66153 | 0.7683 | 0.63904 | 0.58556 | 0.68415 | 0.714 | 0.82736 | 0.77988 | 0.69293 | 0.6381 | 0.72904 | 0.63548 | 0.61923 | 0.62691 | 0.52035 | 0.63826 | 0.68295 | 0.51902 | 0.54348 | 0.58045 | 0.5009 | 0.48708 | 0.421 | 0.38598 | 0.44554 | 0.24486 | 0.40537 | 0.41281 | 0.31471 | 0.23953 | 0.24935 | 0.18492 | 0.24211
-77.13 | 0.24458 | 0.28444 | 0.31452 | 0.26917 | 0.33214 | 0.4437 | 0.38426 | 0.45713 | 0.59944 | 0.68827 | 0.73377 | 0.66242 | 0.89675 | 0.87825 | 0.88348 | 0.77129 | 0.80658 | 0.8352 | 0.81832 | 0.6639 | 0.69017 | 0.74225 | 0.87795 | 0.83501 | 0.98223 | 0.72208 | 0.70015 | 0.65416 | 0.76999 | 0.62074 | 0.65378 | 0.73554 | 0.69903 | 0.62528 | 0.6807 | 0.64398 | 0.59536 | 0.61735 | 0.69709 | 0.58721 | 0.58323 | 0.47024 | 0.4642 | 0.3185 | 0.37085 | 0.34668 | 0.43935 | 0.34345 | 0.42923 | 0.23047 | 0.15428
-77.12 | 0.28329 | 0.2946 | 0.30103 | 0.2164 | 0.40966 | 0.45563 | 0.421 | 0.43681 | 0.55917 | 0.63493 | 0.83559 | 0.76099 | 0.94562 | 1.0097 | 0.85549 | 0.94168 | 0.9402 | 0.92263 | 0.83755 | 0.90194 | 0.91166 | 0.85512 | 0.92834 | 0.95613 | 1.0656 | 0.98333 | 0.88 | 0.86143 | 0.95593 | 0.94714 | 0.75679 | 0.85821 | 0.90732 | 0.80329 | 0.75456 | 0.80196 | 0.80657 | 0.75029 | 0.75439 | 0.66163 | 0.74387 | 0.54666 | 0.53344 | 0.35902 | 0.46043 | 0.37164 | 0.32608 | 0.34352 | 0.24191 | 0.26831 | 0.15701
-77.11 | 0.26907 | 0.27232 | 0.2589 | 0.31733 | 0.38187 | 0.45206 | 0.56769 | 0.44824 | 0.65787 | 0.76504 | 0.95479 | 1.0158 | 1.0393 | 1.1021 | 1.1165 | 1.1154 | 1.1293 | 1.0136 | 0.99754 | 1.006 | 1.1089 | 1.0599 | 0.99965 | 1.0107 | 1.0781 | 1.1075 | 0.97005 | 0.89659 | 0.87196 | 0.90675 | 0.94108 | 0.91007 | 0.95267 | 0.91976 | 0.94801 | 0.89918 | 0.87918 | 0.86711 | 0.82505 | 0.85901 | 0.74481 | 0.68669 | 0.50064 | 0.45253 | 0.45065 | 0.35947 | 0.43013 | 0.36633 | 0.2614 | 0.24784 | 0.16068
-77.1 | 0.22338 | 0.23548 | 0.27834 | 0.29062 | 0.38782 | 0.54648 | 0.49601 | 0.57948 | 0.56588 | 0.72942 | 0.8989 | 0.96002 | 1.2048 | 1.0913 | 1.1794 | 1.1138 | 1.3584 | 1.1859 | 1.3023 | 1.1287 | 1.176 | 1.1514 | 1.2286 | 1.1439 | 1.0989 | 1.1365 | 0.98213 | 0.98677 | 1.1396 | 1.0569 | 1.1605 | 1.0853 | 1.1208 | 1.0105 | 0.96832 | 0.96285 | 1.0183 | 0.99179 | 0.90116 | 0.92334 | 1.02 | 0.83219 | 0.6207 | 0.49692 | 0.43449 | 0.41668 | 0.45487 | 0.3537 | 0.35196 | 0.23313 | 0.17484
-77.09 | 0.25443 | 0.28701 | 0.31561 | 0.36213 | 0.49023 | 0.56324 | 0.57461 | 0.61752 | 0.60889 | 0.71858 | 0.7978 | 1.0083 | 1.139 | 1.1039 | 1.2864 | 1.3283 | 1.4373 | 1.6183 | 1.712 | 1.5651 | 1.5337 | 1.3283 | 1.3742 | 1.2909 | 1.2569 | 1.1724 | 0.9496 | 0.92077 | 1.1159 | 1.1446 | 1.2567 | 1.311 | 1.4286 | 1.4184 | 1.4896 | 1.2593 | 1.2785 | 1.1583 | 1.1166 | 0.98351 | 0.98818 | 0.83779 | 0.71443 | 0.54396 | 0.54678 | 0.58872 | 0.42624 | 0.47214 | 0.27156 | 0.32135 | 0.20558
-77.08 | 0.25769 | 0.23636 | 0.35019 | 0.4294 | 0.5524 | 0.68017 | 0.79844 | 0.65987 | 0.78399 | 0.87738 | 1.0054 | 1.0096 | 1.2821 | 1.2661 | 1.4856 | 1.5034 | 2.0201 | 1.9699 | 2.4111 | 2.3111 | 2.2193 | 1.7592 | 1.69 | 1.4731 | 1.3639 | 1.1888 | 1.0792 | 1.1088 | 1.2007 | 1.3108 | 1.5353 | 1.7364 | 2.0445 | 2.11 | 2.0973 | 1.8897 | 1.6892 | 1.4369 | 1.2643 | 1.2703 | 1.0929 | 0.87756 | 0.78419 | 0.72658 | 0.6301 | 0.54361 | 0.49269 | 0.46094 | 0.40684 | 0.24258 | 0.20017
Pixels 6 & 7
160 m
Good pixels separation good position resolution
Good for particle and X-ray imaging detector
Laser Study of 49x54 3D-Trench Electrode Detector Array
2017 results

## Slide 53
探测器模拟、设计、开发及制作的团队。
The Detector Team at Xiangtan University

## Slide 54
Summary
研究背景
一
They can form SDD array up to 1 m2
Good for X-ray Applications
2
4
1
3D-Terench-Electrode Detectors has been developed
3
Large area SDD is now designed
with Φ up to 5 cm
3D-Terench-Electrode Detectors are good for X-ray and high energy physics applications

## Slide 55
14
Thank You！

## Slide 56
高性能硅探测器
6
三维沟槽电极硅探测器
国际首次提出
基于SOI的高阻硅CMOS MAPs探测器
填补国内空白
超大面积硅漂移室探测器
国际首次提出
噪声低、功耗低、探测效率高
对比技术指标 | 传统探测器 | 高性能探测器
能量分辨率 | ≥200 eV | ~130 eV
工作电压 (漏电流相同) | 100 V | ~4 V
抗辐射效率 | 5% (1016 n/cm2) | ≥25% (1016 n/cm2)
收集时间 | ~5 ns | ~50 ps
面积 | 25 mm2 | 30 cm2

## Slide 57
高性能硅探测器的重大需求
7
医疗设备和医学检测（约10亿元/年）
核工业核探测（约5亿元/年）
新一代同步光源（约5亿元/年）
航天器导航系统（约40亿元/年）
湖南益阳正在筹建的核电站
目前国内高性能探测器全部依赖进口，涉及到国防及空天领域的高性能探测器国外对我国实施严格控制或禁运。高性能探测器的技术空白已经严重制约了我国国民经济的发展，威胁到我国国防安全！

## Slide 58
湘潭大学半导体粒子光子探测器中心 （2015年底建成）

## Slide 59
1
2
2
8.7
7
4
3
7.0
5.7
1.5
7.85
7.85
2
3.5
3.5
3.925
6.5
2
2
2
3
1.5
4
2
2
2
2
2.5
2.5
3
6.5
8.5
8
15
3.925
3.925
3.925
3.0
2
2
1.5
五百级
五百级
办公室 I
办公室 II
办公室 III
办公室 IV
实验室 B
实验室 A
I-V测试仪、C-V测试仪
楼梯
休息室
百级
百级
男 厕
女 厕
小型会议室
高层办公室
大型会议室
李总办公室
洗手间
秘书
办公室
电 梯
激光切片机
晶片存储柜
磁控溅射仪
原子层沉积
圆片烧结炉
金相显微镜
不锈钢工作台
金相显微镜
通风橱
双面光刻机
甩干机
激光直刻机
风淋室
更衣室
通风厨
甩干机
氧化炉
净水系统
晶片工作台
超声波圆片清洗机
职工活动室
职工休息室
楼 梯
低温等离子体化学气相沉积(LPCVD)二氧化硅系统
等离子体增强化学汽相沉积(PECVD)多晶硅系统
深反应离子刻蚀机
通风橱
双面光刻机
百级
五百级
2.5
1
1
3
2.5
2
1
2
2
2
18m
6.3m
8
1
2
玻璃门
1.5
门
窗
N
1.5
1.5
4.0
3.0
3.0
3.8
1.2
3.0
6.3m
shower
坐便器
47.4m
湘潭市高新区超纯高阻硅探测器项目 （2016年6月建成）